## Slide 1: Instrument Qualification
Mohammed Shafikur Rahman, PhD
Assistant Professor

## Slide 2: OBJECTIVES
• Concepts and Principles in HPLC Qualification and Calibration
– Overview of equipment qualification
– Intro to HPLC, a dominant analytical tool for research and quality control
– HPLC equipment qualification vs. calibration
– DQ, IQ, OQ and PQ
• A Procedure for HPLC Calibration and Operational Qualification
– Outsourcing vs. internal execution
– Describe a generic HPLC calibration procedure (OQ)
– Setting acceptance criteria for pump, detector, autosampler and column oven
– Learn how to calibrate an HPLC system within one day using internal personnel

## Slide 3: Instrument/Equipment
Equipment is the collective analytical measurement instruments assembled to perform a mechanical Process.
Equipment qualification is the action of proving that any equipment works correctly and actually leads to accurate and reliable results.
Qualification is an act or process to assure something complies with some condition, standard, or specific requirements. It applies on one element of process to be validated.
Validation:proof document that the process will consistently produce a predetermine outcome

## Slide 4: Phase of validation
User Requirement Specification (URS)
Design qualification (DQ)
Factory Acceptance Test (FAT)
Installation Qualification (IQ)
Site Acceptance Test (SAT)
Operation Qualification (OQ)
Performance Qualification (PQ)
Maintenance Qualification (MQ)
Calibration/Component Qualification (CQ)

## Slide 5: Instrument Qualification
Time line
Pre-installation Upon Installation at User Site Post-Installation
URS+DQ+FAT SAT+IQ OQ PQ Calibration and maintenance
Defining URS and setting functional and performance specification
Verification procedure and documentation of the installation
Equipment testing to verify system is operating to predefined functional and performance specification
Perform periodic performance verification (OQ) and preventive maintenance
Verification that system is consistently operating as intended for the selected application. (system suitability testing
Initially Ongoing

## Slide 8: URS
User Requirements Specification (URS), is the most critical of documents and yet, the most often bungled. Whether the system is purely mechanical, or a mix of electro-mechanical, or solely a software program, the successful compilation and execution of the Installation Qualification (IQ) (for installation), Operational Qualification (OQ) (for functionality) and the Performance / Product Qualification (PQ) (for operability), is dependent on an User Requirements Specification (URS) containing clear, concise and testable requirements.

## Slide 9: URS cont…
The following guidelines should be followed during the production of the URS :
1. Each requirement statement to be uniquely referenced,and no longer than 250 words.
2. Requirement statements should not be duplicated nor contradicted.
3. The URS should express requirements and not design solutions.
4. Each requirement should be testable.

## Slide 10: URS cont…
5. The URS must be understood by both user and supplier; ambiguity and jargon should be avoided.
6. The use of diagrams is often useful.
7. The scope for readers to make assumptions or misinterpret should be minimized.
8. Wherever possible, the URS should distinguish between mandatory/regulatory requirements and desirable features.

## Slide 11: Validation vs Qualification
Validation:
Refers to the total life cycle of a product from development through use and maintenance.  Owners are responsible for Validating Their Processes (personnel, equipment, methods, SOPs) to ensure compliance to cGMP/GLP regulations.
Qualification: (Inspection, functional testing and documentation review)
Is a part of the validation process which verifies module and system functional performance prior to being placed on-line and thereafter according to a standard operating procedure.

## Slide 12: Equipment Qualification
It is a basic requirement of good analytical chemistry that balances and other analytical instruments must be suitable for the purpose for which they are used and that they must be appropriately calibrated. As a consequence, Equipment Qualification is gaining more and more importance in ensuring the validity of results. Regulatory bodies also seem to be turning their attention increasingly to this area, and manufacturers of analytical equipment are forced to play a significant role in the various steps of Equipment

## Slide 13: 5 steps of Equipment Qualification
Step 1: Design Qualification (DQ) defines the functional and operational specifications of a balance or instrument.
Step 2: Installation Qualification (IQ) ensures that a balance or instrument is received as designed and specified. It documents the installation in the selected user environment.
Step 3: Operational Qualification (OQ) demonstrates that a balance or instrument will function according to its operational specification in the selected environment.
Step 4: Performance Qualification (PQ) demonstrates that a balance or instrument consistently performs according to a specification appropriate to its routine use.
Step 5: Maintenance Qualification (MQ) describes and documents any maintenance required on the equipment.

## Slide 14: DQ and vendor qualification
As part of the design qualification process, the vendor should be qualified; there may be situations where a vendor audit is recommended.

## Slide 15: IQ includes
• A procedure designed to establish that the instrument is installed and configured according to the requirements established by the instrument manufacturer.
• Instrument functional description
– Specific instrument ID (i.e. serial number), location
– Software / firmware revision
– Verification of certificates and manuals • Installation/configuration procedures
– Site compatibility
– Electrical, fluidic, communication connections • User approval of IQ completion
– With signature and date

## Slide 16: OQ typically includes
• A procedure designed to establish that the system is functioning according to manufacturer specifications throughout its intended operating range.
• Functional qualification procedure
– Power-up initialization and diagnostic for each LC component
• Acceptance criteria
• User’s approval of OQ completion
-with signature and date

## Slide 17: What is PQ?
• A procedure designed to establish that the overall system is performing as intended with actual use.
– If electronic data is stored via a computer network, network access should be verified as part of the performance qualification.
• PQ protocols are developed by the user, though generic procedures using test mixtures might be adequate
• Typically, PQ involves performing system suitability test.
– Check precision, peak resolution, tailing factor, etc

## Slide 18: Maintenance Qualification
The MQ describes and documents any maintenance required on the equipment. This includes routine servicing and any repairs necessary. Details of any maintenance contracts are also documented in this section, together with a list of authorized service engineers. In addition, the MQ includes the routine cleaning of the equipment and also its ultimate disposal.

## Slide 19: Component Qualification/Calibration
Component qualification (CQ) – is a relatively new term developed in 2005. This term refers to the manufacturing of auxiliary components to ensure that they are manufactured to the correct design criteria. This could include packaging components such as folding cartons, shipping cases, labels or even phase change material. All of these components must have some type of random inspection to ensure that the third party manufacturer's process is consistently producing components that are used in the world of GMP at drug or biologic manufacture.

## Slide 21: Instrument re-qualification
Instrument Validation should not be viewed as a one-off event – confidence in analytical results is required for the whole of the instrument’s working life.  To ensure that this confidence is retained, the instrument validation process should be repeated at regular intervals during the instruments operational life.  The difference between Installation Validation and Re Qualification is that IQ is omitted for the Re-Qualification  Re-Qualification should be performed at least annually and should be performed more frequently for applications whose test results have critical im